## Sheet
FUTURA-System-Manual
Table 1, Page 2
Version
FuturaTool | 3.0
FUTURA Firmware | 2330-07-125
V350 Firmware | 35010710
Generic product label identification
Table 2, Page 13
Symbol | Meaning
This product complies with the essential EEA requirements for the Electromagnetic Compatibility (EMC) Directive 2014/30/EU, the Low Voltage Directive (LVD) 2014/35/EU and the Restriction of the use of certain Hazardous Substances Directive (RoHS) 2011/65/EU
This product also complies with the following Statutory Instruments. For a list of standards, refer to our UKCA DoC or contact Aber on the above for EU DoC. SI 2016 No. 1101 - The Electrical Equipment (Safety) Regulations 2016*. SI 2016 No. 1091 - The Electromagnetic Compatibility Regulations 2016*. SI 2012 No. 3032 - The Restriction of the Use of Certain Hazardous Substances in Electrical and Electronic Equipment Regulations 2012*. *As amended by EU Exit Regulations.
Contains hazardous components and must not be disposed of at a household waste site. Instead it should be taken to the appropriate collection point for the recycling of waste electrical and electronic equipment (WEEE).
Power supply ratings are 24 V DC at 2.5 A. Negative is connected to the outer contact and the centre pin is connected to the positive supply on this terminal.
Continued on next page
Table 3, Page 14
Symbol | Meaning
Use only Aber Instruments supplied Power Supply Unit SPU61A-108-ABER with this product.
USB
Refer to User Manual
Keep dry and away from moisture
Indoor use only
Caution, risk of electric shock
Caution, risk of danger
(NRTL) UL/CSA 61010-1 (3rd Ed.) Am. 1, Electrical Equipment for Measurement, Control, and Laboratory Use; Part1:General Requirements, Rev Nov 2018
Alternating current
Ground symbol
Table 4, Page 20
Part Number | Description | Overall dimensions (mm)
|  | W | X | Y | Z
6531-52 | Annular Electrodes 12 x 120mm Stainless Steel / Fortron | 105 | 223 | 152 | -
6531-52/PG | Pre-fitted PG13.5 nut. Annular Electrodes 12 x 120mm Stainless Steel / Fortron | - | 223 | 152 | 128
6532-52 | Annular Electrodes 12 x 220mm Stainless Steel / Fortron | 205 | 323 | 252 | -
6532-52/PG | Pre-fitted PG13.5 nut. Annular Electrodes 12 x 220mm Stainless Steel / Fortron | - | 323 | 252 | 228
6530-52 | Annular Electrodes 12 x 320mm Stainless Steel / Fortron | 305 | 423 | 352
6530-52/PG | Pre-fitted PG13.5 nut. Annular Electrodes 12 x 320mm Stainless Steel / Fortron | - | 423 | 352 | 328
6534-52 | Annular Electrodes 12 x 450mm Stainless Steel / Fortron | 435 | 555 | 482 | -
6534-52/PG | Pre-fitted PG13.5 nut. Annular Electrodes 12 x 450mm Stainless Steel / Fortron | - | 555 | 482 | 458
6570-52 | Pre-fitted PG13.5 nut. Annular Electrodes 12 x 320mm Stainless Steel / Fortron (Kleenpak) | - | 423 | 352 | 285
Table 5, Page 21
Part Number | Description | Overall dimensions (mm)
|  | W | X | Y | Z
6543-52 | Flush Electrodes 25 x 175mm Stainless Steel / Fortron | 25 | 175 | 239 | 22
6541-52 | Flush Electrodes 25 x 75mm Stainless Steel / Fortron standard O-ring position | 25 | 75 | 132 | 22
6542-52 | Flush Electrodes 25 x 75mm Stainless Steel / Fortron for use with some Braun safety ports | 25 | 75 | 132 | 46
6545-52 | Annular Electrodes 25 x 96mm Stainless Steel / Fortron standard O-ring position | 25 | 96 | 154 | 22
6546-52 | Annular Electrodes 25 x 96mm Stainless Steel / Fortron for use with some Braun safety ports | 25 | 96 | 154 | 47
Table 6, Page 22
Part Number | Description | Overall dimensions (mm)
|  | W | X | Y | Z
6551-52 | Probe Flush Electrodes 25 x 175mm Stainless Steel / Fortron with size 1 – 1.5” Tri-clamp flange, can be used with NovAseptic ports | 25 | 50.4 | 175 | 225
6552-52 | Probe Flush Electrodes 25 x 70 mm Stainless Steel / Fortron with size 1 – 1.5” Tri-clamp flange, can be used with NovAseptic ports | 25 | 50.4 | 70 | 120
6553-52 | Probe Flush Electrodes 25 x 70 mm Stainless Steel / Fortron with size 2” Tri-clamp flange, can be used with NovAseptic ports | 25 | 63.9 | 70 | 120
Table 7, Page 23
Part Number | Description | Overall dimensions (mm)
|  | X | Y | Z
5320-00 | Pico Electrodes 7.5 x 140mm Vectra MT4310 | 190 | 140 | 7.5
All values assume use with a standard range probe. Values may differ if using an extended range probe.
Table 8, Page 24
Product Name: | Standard FUTURA | Standard Remote FUTURA | FUTURA Pico
Frequency Range | 50 kHz to 20 MHz
Measuring Ranges | Capacitance: 0.0 to 400 pF/cm Conductivity: 1.0 to 40 mS/cm +/-0.1 mS/cm (Higher ranges available with compatible probes)
Cell Concentration Range | Depends on cell sizes but typically: μ 6 10 Yeast (6 m): 10 cells/ml to 10 cells/ml μ 9 13 Bacteria (1 m): 10 cells/ml to 10 cells/ml μ 5 9 Animal Cell (12 m): 10 cells/ml to 10 cells/ml μ 3 7 Plant Cell (50 m): 10 cells/ml to 10 cells/ml
Resolution | Resolution: 0.1 pF/cm. Approx: 9 Bacteria typically 2 g/L dry weight or 2x10 cells/ml for E. Coli. 5 Yeast or Animal Cells 0.05 g/L or 1x10 cells/ml The relationship of these capacitance values to biomass levels depend upon the cell type and line.
Accuracy | Typically better than ±3 % across the measurement range or ±2 % of the reading
Stability | Better than ±0.2 pF/cm at constant temperature with standard conductivity solution ∼ of 12 mS/cm
Linearity | Better than ±1 % over 100 pF/cm
Precision | Typically <±0.5 pF/cm, no filter active
Power Supply | Natively 24 V DC
Environmental | IP65 rated |  | 1 IP6* rated
Main Enclosure Dimensions (approx):
Height | 50 mm | 30 mm | 27 mm
Width | 130 mm | 135 mm | 85 mm
Depth | 64 mm | 64 mm | 50 mm
Weight | 375 g | 211 g | 225 g
Table 9, Page 25
Standard Remote FUTURA: Remote Enclosure Dimensions (approx)
Height | 135 mm
Diameter | 24 mm
Weight | 203 g
2.3.1 FUTURA Connect Specifications
Table 10, Page 28
Product Name | FUTURA Connect 1 | FUTURA Connect 4
Part Number | 2801-00 | 2814-00
General | 1-way FUTURA Hub 1 FUTURA to 1 USB, 1 Modbus & 2 Active Current Loop connections | 4-way FUTURA Hub 4 FUTURAs to 1 USB, 1 Modbus & 8 Active Current Loop connections
Communications | 1 x FUTURA Connection | 4 x FUTURA Connections
1 x USB 2.0
1 x Modbus RTU
| 2 x 4-20 mA Current Loops (2 per FUTURA) | 8 x 4-20 mA Current Loops (2 per FUTURA)
Power Supply | Desktop Mains PSU supplied by Aber. (Part number: 2885-35, AC Input 110 to 240 volts AC 50/60 Hz). Installation Category II
Power Consumption | 10 W Maximum | 60 W Maximum
Electrical Safety | For SAFETY reasons must ONLY use Aber external power supply, part number 2885-35. Only AC mains power cord supplied by Aber Instruments may be used with the Desktop mains power supply.
Enclosure Dimensions (approx) Height Width Depth Weight
| 38 mm | 88 mm
| 105 mm | 175 mm
| 145 mm | 132 mm
| 0.7 kg | 1.35 kg
Environmental | Safe ambient operating temperature range: 5 °C to 40 °C Recommended ambient operating temperature range: 15 °C to 30 °C Relative Humidity: <85 %. Pollution: Deg 2. (EN61010) Not Waterproof. Recommended operating altitude: <2000 m
FUTURA Connect Xtra Specifications
Table 11, Page 29
Product Name | FUTURA Connect Xtra
Part Number | 2815-00
General | 1-way FUTURA Hub 1 FUTURA to 1 USB, 1 Modbus connection
Communications | 1 x FUTURA Connection
1 x USB 2.0
1 x Modbus RTU via Anybus plug-in module (See 7.4.1 The Connect Xtra & Available network capabilities)
Power Supply | Desktop Mains PSU supplied by Aber. (Part number: 2885-35, AC Input 110 to 240 volts AC 50/60 Hz). Installation Category II
Power Consumption | 10 W Maximum
Electrical Safety | For SAFETY reasons must ONLY use Aber external power supply, part number 2885-35. Only AC mains power cord supplied by Aber Instruments may be used with the Desktop mains power supply.
Enclosure Dimensions (approx) Height Width Depth Weight
38 mm
105 mm
145 mm
0.7 kg
Environmental | Safe ambient operating temperature range: 5 °C to 40 °C Recommended ambient operating temperature range: 15 °C to 30 °C Relative Humidity: <85 %. Pollution: Deg 2. (EN61010) Not Waterproof. Recommended operating altitude: < 2000m
2.4.1 V350 Touch - Panel Mount Specifications
Table 12, Page 30
Part Number | 3501-00
General | Touch Screen Interface Two 4 to 20 mA Current Loops V350 Connection Lead, Aber Part # 3512-11
Measuring Ranges | Capacitance: 0 to 400 pF/cm, equivalent cells/ml Conductivity: 1 to 40 mS/cm.
Communications | 1 x FUTURA pass through USB connection 1 x FUTURA Connection 2 x 4 to 20 mA Active Current Loops 1 x Alarm Relay
Power Supply | +24 V DC
Power Consumption | 15 W Maximum
Electrical Safety | Must be mounted in a suitable EMC shielded enclosure & grounded. Refer to installation instructions.
Battery Backup | 7 years typical, at 25 °C
Overall Dimensions (approx)
Height | 114 mm
Width | 125 mm
Depth | 87 mm (allow for power supply cables)
Weight | 0.5 kg
Environmental | IP65 for fascia, when mounted in a suitable panel. Safe ambient operating temperature range: 5 °C to 40 °C Recommended ambient operating temperature range: 15 °C to 30 °C Relative Humidity: < 85 %. Pollution: Deg 2. (EN61010) Recommended operating altitude: < 2000 m
Disposal | This equipment must be disposed of in accordance with the Waste from Electrical and Electronic Equipment (WEEE) Directive - European Standard EN 50419. This product must not be treated as household waste. Instead it shall be handed over to an appropriate collection point for the recycling of electrical and electronic equipment. If in doubt, please return this equipment to Aber Instruments Ltd. who will dispose of it correctly for you.
2.4.3 V350 Touch - Wall Mount Specifications
Table 13, Page 32
Part Number | 3550-00
General | Touch Screen Interface Two 4 to 20 mA Current Loops Mounted in Stainless Steel Wall Mountable Box
Measuring Ranges | Capacitance: 0 to 400 pF/cm, equivalent cells/ml Conductivity: 1 to 40 mS/cm.
Communications | 1 x FUTURA pass through USB connection 1 x FUTURA Connection 2 x 4 to 20 mA Active Current Loops 1 x Alarm Relay
Power Supply | +24 V DC
Power Consumption | 15 W Maximum
Electrical Safety | Must be mounted in a suitable EMC shielded enclosure & grounded. Refer to installation instructions.
Battery Backup | 7 years typical, at 25 °C
Overall Dimensions (approx)
Height | 200 mm (allow for cable entry)
Width | 165 mm
Depth | 135 mm (max)
Weight | 2.8 kg
Environmental | IP65 for fascia, when mounted in a suitable panel. Safe ambient operating temperature range: 5 °C to 40 °C Recommended ambient operating temperature range: 15 °C to 30 °C Relative Humidity: < 85 %. Pollution: Deg 2. (EN61010) Recommended operating altitude: < 2000 m
Disposal | This equipment must be disposed of in accordance with the Waste from Electrical and Electronic Equipment (WEEE) Directive - European Standard EN 50419. This product must not be treated as household waste. Instead it shall be handed over to an appropriate collection point for the recycling of electrical and electronic equipment. If in doubt, please return this equipment to Aber Instruments Ltd. who will dispose of it correctly for you.
Table 14, Page 51
Signal Simulator |  | FUTURA Readings
Capacitance | Conductivity | Capacitance pF/cm | Conductivity mS/cm
100 | 40 | 100 ±3 | 40 ±1
0 | 0 | 0 ±1 | 0 ±0.5
Table 15, Page 52
Minimum probe clearance | Standard | Flush | Pico
A : Immersion depth (mm) | 60 | 40 | 30
B : Side clearance (mm) | 40 | 40 | 20
C : Base clearance (mm) | 40 | 40 | 20
Table 16, Page 53
KCl solution 12.88 mS/cm | Criteria | Corrected Conductivity | Measured values
Capacitance | 7 pF/cm ±2 pF/cm | - | pF/cm
Conductivity | 12.88 mS/cm ±0.2 mS/cm @25 °C See Note 1 |  | mS/cm
Temperature | °C | - | °C
values (check the
Table 17, Page 53
22 °C | 12.15 mS/cm
23 °C | 12.39 mS/cm
24 °C | 12.64 mS/cm
25 °C | 12.88 mS/cm
26 °C | 13.13 mS/cm
27 °C | 13.37 mS/cm
Use the table format overleaf to record your readings.
Table 18, Page 54
Step | Sample | Capacitance pF/cm | Conductivity mS/cm | Comments
1 | Medium / KCl
2 | Cell Sample
3 | 50:50
4
5
6
7
8
9
10
Table 19, Page 105
V350 Touch - Wall Mount 14 Way Terminal Strip
Terminal | Function |  | Terminal | Function
Active Current Loops
1 | 0 V DC |  | 8 | 0 V
2 | 24 V DC |  | 9 | Capacitance 4 to 20 mA
|  |  | 10 | 0 V
3 | Reserved, do not use |  | 11 | Conductance 4 to 20 mA
4
5 |  |  |  | Alarm Outputs
6 |  |  | 12 | Alarm Relay Common
7 |  |  | 13 | Alarm Relay Switched
|  |  | 14 | Reserved, do not use
values):
Table 20, Page 138
22 °C | 12.15 mS/cm
23 °C | 12.39 mS/cm
24 °C | 12.64 mS/cm
25 °C | 12.88 mS/cm
26 °C | 13.13 mS/cm
27 °C | 13.37 mS/cm
Device, Time and Measurements
Table 21, Page 150
Device Name | Device Serial | Probe Serial
My FUTURA | HARSc15478 | TCDK11396
Time Stamp | Run Time (Decimal Hours) | Biomass | Capacitance (pF/cm) | Conductivity (mS/cm)
Measurement Settings
Table 22, Page 150
Measuring Frequency (kHz) | Filter | PoIC | Dual Frequency | High Frequency (kHz)
Secondary Measurements and Events
Table 23, Page 150
Capacitance @ Measuring Frequency (pF/cm) | Capacitance @ High Frequency (pF/cm) | Events | Probe Temperature (°C)
Device, Time and Events
Table 24, Page 151
Device Name | Device Serial | Probe Serial
My FUTURA | HARSc15478 | TCDK11396
Time Stamp | Run Time (Decimal Hours) | Events
Measurements
Table 25, Page 151
Capacitance @ 50 kHz (pf/cm) | ... | Capacitance @ 20000 kHz (pf/cm) | Conductivity @ 50 kHz (mS/cm) |  | Conductivity @ 20000 kHz (mS/cm)
Settings and Temperature
Table 26, Page 151
Filter | PoIC | Probe Temperature (°C)
List of frequencies measured when
Table 27, Page 153
Reading number | Frequency
1 | 50 kHz
2 | 64 kHz
3 | 82 kHz
4 | 106 kHz
5 | 136 kHz
6 | 174 kHz
7 | 224 kHz
8 | 287 kHz
9 | 368 kHz
10 | 473 kHz
11 | 580 kHz
12 | 779 kHz
13 | 1000 kHz
14 | 1120 kHz
15 | 1648 kHz
16 | 2115 kHz
17 | 2714 kHz
18 | 3484 kHz
19 | 4472 kHz
20 | 5740 kHz
21 | 7368 kHz
22 | 9457 kHz
23 | 12139 kHz
24 | 15650 kHz
25 | 20000 kHz
Table 28, Page 169
Name | Length | Function
Start | 3.5 x character duration | Indicates start of message
Address | 1 byte | Modbus address of receiving device
Function | 1 byte | Type of action to perform. See Memory Map section below
Data | 2 bytes | The actual data being transmitted
CRC Check | 2 bytes | Error checking for data integrity
End | 3.5 x character duration | Indicates end of message
below:
Table 29, Page 170
Modbus Command | Description
0x03 | Read Holding Registers
0x04 | Read Input Registers
0x06 | Write Single Register
0x08 | Diagnostics
0x10 | Write Multiple Registers
0x0E / 0x2B | Read Device Identification
Note: 32-bit floats in the table below are IEEE754 format 32-bit floating point values.
Table 30, Page 171
Address | Name | Permission | Values and Scaling | Remember?
0 | Measurement Mode | Read + Write | 0 = Cell Culture 1 = Bacterial 2 = Microbial 3 = Signal Simulator 4 = Custom 5= Brewery 6=Scan | Yes
1 | FUTURA Status Flags | Read | 0 – 65535 (see following table)
2 | Baud Rate | Read | 0 – 7 (see following table) | Yes
3 | Parity | Read | 0 – 2 (see following table) | Yes
4 | Modbus Address | Read | 2 | Yes
5 | Capacitance Measurement (32 bit float – msb) | Read | Units: pF/cm Scaling: top 2 bytes of 4 byte floating point number
6 | Capacitance Measurement (32 bit float – lsb) | Read | Scaling: lower 2 bytes of 4 byte floating point number
7 | Conductance Measurement (32 bit float – msb) | Read | Units: mS/cm Scaling: top 2 bytes of 4 byte floating point number
8 | Conductance Measurement (32 bit float – lsb) | Read | Scaling: lower 2 bytes of 4 byte floating point number
Table 31, Page 172
Address | Name | Permission | Values and Scaling | Remember?
9 | Measuring Frequency (16 bit unsigned integer) | Read + Write | Range: 50 to 20000 Units: kHz | Yes
10 | High Frequency (Dual) (16 bit unsigned integer) | Read + Write | Range: 1000 to 20000 Units: kHz | Yes
11 | Polarisation Correction | Read + Write | 0 = No Pol C 1 = 220 Pol C 2 = FUTURA Pol C | Yes
12 | Filter Level | Read + Write | Range: 0 to 200 0 and 1 = Off > 2 = On, value is length of running average filter. | Yes
13 | Custom Calibration | Read + Write | 0 = Use factory TIP cal. 1 = Use custom cal 1. 2 = Use custom cal 2. | Yes
14 | Zero Value (32 bit float – msb) | Read + Write | Shows the Zero offset. Units: pF/cm Scaling: top 2 bytes of 4 byte floating point number (Write 0.0 to registers 14-15 to clear a ‘single’ frequency zero) | Yes
15 | Zero Value (32 bit float – lsb) | Read + Write | Scaling: lower 2 bytes of 4 byte floating point number | Yes
16 | Clean Pulse Now | Read + Write | A change from 0 to 1 triggers a clean pulse function.
17 | Zero Now | Read + Write | A change from 0 to 1 triggers a zero function. Writing a value of 2 will clear a frequency scan zero but won’t clear a ‘single’ frequency zero.
18 | Biomass Multiplying Factor CPM (32 bit float – msb) | Read + Write | Units: user defined (cells/ml) Scaling: top 2 bytes of 4 byte floating point number | Yes
19 | Biomass Multiplying Factor CPM (32 bit float – lsb) | Read + Write | Scaling: lower 2 bytes of 4 byte floating point number | Yes
20 | Biomass Measurement CPM (32 bit float – msb) | Read | Units: user defined (cells/ml) Scaling: top 2 bytes of 4 byte floating point number
21 | Biomass Measurement CPM (32 bit float – lsb) | Read | Scaling: lower 2 bytes of 4 byte floating point number
Table 32, Page 173
Address | Name | Permission | Values and Scaling | Remember?
22 | Probe Temperature Measurement (32 bit float – msb) | Read | Units: ˚C Scaling: top 2 bytes of 4 byte floating point number
23 | Probe Temperature Measurement (32 bit float – lsb) | Read | Scaling: lower 2 bytes of 4 byte floating point number
24 | Probe Total Hours Use (32 bit un int – msb) | Read | Units: Hours Scaling: top 2 bytes of 4 byte unsigned integer
25 | Probe Total Hours Use (32 bit un int – lsb) | Read | Scaling: lower 2 bytes of 4 byte unsigned integer
26 | Probe Total CIP Autoclaves Experienced (32 bit un int – msb) | Read | Scaling: top 2 bytes of 4 byte unsigned integer
27 | Probe Total CIP Autoclaves Experienced (32 bit un int – lsb) | Read | Scaling: lower 2 bytes of 4 byte unsigned integer
28 | Probe CIP & Autoclave Threshold (32 bit un int – msb) | Read | Scaling: top 2 bytes of 4 byte unsigned integer
29 | Probe CIP & Autoclave Threshold (32 bit un int – lsb) | Read | Scaling: lower 2 bytes of 4 byte unsigned integer
30 | Probe Leakage Detections (32 bit un int – msb) | Read | Scaling: top 2 bytes of 4 byte unsigned integer
31 | Probe Leakage Detections (32 bit un int – lsb) | Read | Scaling: lower 2 bytes of 4 byte unsigned integer
32 | Probe Leakage Detections Threshold (32 bit un int – msb) | Read | Scaling: top 2 bytes of 4 byte unsigned integer
33 | Probe Leakage Detections Threshold (32 bit un int – lsb) | Read | Scaling: lower 2 bytes of 4 byte unsigned integer
34 | Increment Autoclave Count Manually | Read + Write | A change from 0 to 1 triggers an increase of 1 to the Autoclave count of the attached probe.
Table 33, Page 174
Address | Name | Permission | Values and Scaling | Remember?
35 | FUTURA Serial Number (Characters 1 + 2) | Read | The two bytes are individual ASCII characters.
36 | FUTURA Serial Number (Characters 3 + 4) | Read | The two bytes are individual ASCII characters.
37 | FUTURA Serial Number (Characters 5 + 6) | Read | The two bytes are individual ASCII characters.
38 | FUTURA Serial Number (Characters 7 + 8) | Read | The two bytes are individual ASCII characters.
39 | FUTURA Serial Number (Characters 9 + 10) | Read | The two bytes are individual ASCII characters.
40 | FUTURA Serial Number (Characters 11 + 12) | Read | The two bytes are individual ASCII characters.
41 | FUTURA Serial Number (Characters 13 + 14) | Read | The two bytes are individual ASCII characters.
42 | FUTURA Serial Number (Characters 15 + 16) | Read | The two bytes are individual ASCII characters.
43 | Probe Serial Number (Characters 1 + 2) | Read | The two bytes are individual ASCII characters.
44 | Probe Serial Number (Characters 3 + 4) | Read | The two bytes are individual ASCII characters.
45 | Probe Serial Number (Characters 5 + 6) | Read | The two bytes are individual ASCII characters.
46 | Probe Serial Number (Characters 7 + 8) | Read | The two bytes are individual ASCII characters.
47 | Probe Serial Number (Characters 9 + 10) | Read | The two bytes are individual ASCII characters.
48 | Probe Serial Number (Characters 11 + 12) | Read | The two bytes are individual ASCII characters.
49 | Probe Serial Number (Characters 13 + 14) | Read | The two bytes are individual ASCII characters.
50 | Probe Serial Number (Characters 15 + 16) | Read | The two bytes are individual ASCII characters.
Table 34, Page 175
Address | Name | Permission | Values and Scaling | Remember?
51 | Rocker Filter | N/A | N/A
52 | Firmware ver. (Characters 1 + 2) | Read | The two bytes are individual ASCII characters.
53 | Firmware ver. (Characters 3+ 4) | Read | The two bytes are individual ASCII characters.
54 | Firmware ver. (Characters 5 + 6) | Read | The two bytes are individual ASCII characters.
55 | Firmware ver. (Characters 7 + 8) | Read | The two bytes are individual ASCII characters.
56 | Firmware ver. (Characters 9 + 10) | Read | The two bytes are individual ASCII characters.
57 | Firmware ver. (Characters 11 + 12) | Read | The two bytes are individual ASCII characters.
58 | Firmware ver. (Characters 13 + 14) | Read | The two bytes are individual ASCII characters.
59 | Firmware ver. (Characters 15 + 16) | Read | The two bytes are individual ASCII characters.
60 | Boot Loader ver. (Characters 1 + 2) | Read | The two bytes are individual ASCII characters.
61 | Boot Loader ver. (Characters 3+ 4) | Read | The two bytes are individual ASCII characters.
62 | Boot Loader ver. (Characters 5 + 6) | Read | The two bytes are individual ASCII characters.
63 | Boot Loader ver. (Characters 7 + 8) | Read | The two bytes are individual ASCII characters.
64 | Boot Loader ver. (Characters 9 + 10) | Read | The two bytes are individual ASCII characters.
65 | Boot Loader ver. (Characters 11 + 12) | Read | The two bytes are individual ASCII characters.
66 | Boot Loader ver. (Characters 13 + 14) | Read | The two bytes are individual ASCII characters.
Table 35, Page 176
Address | Name | Permission | Values and Scaling | Remember?
67 | Boot Loader ver. (Characters 15 + 16) | Read | The two bytes are individual ASCII characters.
68 | FUTURA Temperature Measurement (32 bit float – msb) | Read | Units: ˚C Scaling: top 2 bytes of 4 byte floating point number
69 | FUTURA Temperature Measurement (32 bit float – lsb) | Read | Scaling: lower 2 bytes of 4 byte floating point number
70 | Post Clean Pulse blanking period (16-bit unsigned integer) | Read + Write | Range: 0 – 50 minutes | Yes
71 | Unused
72 | Unused
73 | Unused
74 | Unused
75 | F[0] | Read | Scan Frequency 50 kHz
76 | F[1] | Read | Scan Frequency 64 kHz
77 | F[2] | Read | Scan Frequency 82 kHz
: | : | : | :
: | : | : | :
99 | F[24] | Read | Scan Frequency 20,000 kHz
100-101 | C[0] | Read | Capacitance, float
102-103 | G[0] | Read | Conductivity, float
104-105 | C[1] | Read | Capacitance, float
106-107 | G[1] | Read | Conductivity, float
: | : | : | :
192-193 | C[23] | Read | Capacitance, float
194-195 | G[23] | Read | Conductivity, float
196-197 | C[24] | Read | Capacitance, float
198-199 | G[24] | Read | Conductivity, float
Register (Address 1):
Table 36, Page 177
Bit | Symptom | Warning/Error | Bit Mask
0 | Capacitance out of Range | Warning | 0x0001
1 | Conductivity out of Range | Warning | 0x0002
2 | Temperature out of Range | Warning | 0x0004
3 | Probe Leakage | Warning | 0x0008
4 | Clean Leakage | Warning | 0x0010
5 | Reserved |  | 0x0020
6 | FUTURA Memory, File Missing | Error | 0x0040
7 | Probe Memory, File Missing | Error | 0x0080
8 | Reserved |  | 0x0100
9 | No Probe | Indicator | 0x0200
10 | Sig Simulator Connected | Indicator | 0x0400
11-15 | Reserved
Baud Rate Register (Address 2):
Table 37, Page 177
Register value | Baud Rate
0 | 1200
1 | 2400
2 | 4800
3 | 9600
4 | 19200
5 | 38400
Parity Register (Address 3):
Table 38, Page 177
Register value | Parity
0 | None
1 | Odd
2 | Even
Table 39, Page 192
Version | Date | Description
00 | 23/12/2020 | Part number changed to recognise revised Aber template. Incorporated ECN20027 and ECN20028. Removed Connect 8. Updated document styling. TW
01 | 10/02/2021 | Updated date and corrected pictures and formatting, PT
02 | 11/08/2021 | Removed V350 Desktop, updated EMC statement and SLA. Corrected Connect 1 Modbus pinout. TW
03 | 08/11/2021 09/04/2022 | Changed Connect Xtra wattage to 10W, Added Kleenpak probe dimensions PT Added Canadian ICES statement and Clean Pulse blanking period section. TW
04 | 26/04/2022 | Updated EN standard, Probe tester added to 5.5, updated page 32 V350 panel mount Added 7.4.1 Connect Xtra protocol information. Updated V350 sections to widen capacitance, updated firmware version to 233007124 PT
05 | 03/05/2023 | Updated Copyright, Added O Ring and PTFE washers statement, Added Connect Extra Profibus DPV1, KS
06 | 10/01/2024 | Updated Modbus Map to reflect clear frequency scan zero feature. Updated Trademarks, Updated Electromagnetic Compatibility to reflect updated standards, Updated Copyright year, KS